xy - Calibration of 1461J scanner (Peter Martinsson 8/11 2002)

cxx

cyy

dfx

dfy

dsx

dsy

dxx

dyy

sfx

sfy

ssx

ssy

New parmeters (from the calibration):

Old			New			New (after calc.)

X fast sens		381.9 nm/V		372.2 nm/V		374.7 nm/V

X fast derate		0.3649 nm/V2		0.3341 nm/V2		0.3367 nm/V2

X slow sens		435.6 nm/V		418.1 nm/V		420.9 nm/V

X slow derate		0.4600 nm/V2		0.4010 nm/V2		0.4213 nm/V2

Xs-Xf coupling	0.8298 nm/V2		0.7109 nm/V2

Xs-Xf coup der	0.2993 pm/V3		0.2134 pm/V3

Xs-Yf coupling	0.1079 pm/V2		0.1293 pm/V2

Xs-Yf coup der	-0.09777 pm/V3	0.01035 pm/V3

Old			New			New (after calc.)

Y fast sens		377.8 nm/V		368.2 nm/V		369.8 nm/V

Y fast derate		0.3605 nm/V2		0.3282 nm/V2		0.3118 nm/V2

Y slow sens		431.6 nm/V		413.0 nm/V		413.8 nm/V

Y slow derate		0.4650 nm/V2		0.4025 nm/V2		0.4025 nm/V2

Ys-Yf coupling	0.8420 nm/V2		0.7012 nm/V2

Ys-Yf coup der	0.2469 pm/V3		0.05689 pm/V3

Ys-Xf coupling	0.1530 pm/V2		0.2235 pm/V2

Ys-Xf coup der	0.01950 pm/V3	0.2117 pm/V3