Parameters E scanner 4645EV

Calibration june 2003 (Hilde)

THE PARAMETERS FROM THE OLD FILE (AFM DOOR 4645EV (not cal)) GAVE THE BEST VALUES. THESE ARE THE ONES FILLED IN AS CALI-NEW FOR BOTH DOOR AND WINDOW AFM.

Cali-Old			Cali-New

X fast sens		32.22 nm/V			33.29 nm/V

X fast derate		0.02180 nm/V2		0.02482 nm/V2

X slow sens		35.67 nm/V			36.8 nm/V

X slow derate		0.02721 nm/V2		0.03027 nm/V2

Xs-Xf coupling	0.03552 nm/V2		0.05582 nm/V2

Xs-Xf coup der	-0.06251 pm/V3		0.02966 pm/V3

Xs-Yf coupling	0.008010 nm/V2		0.005755 nm/V2

Xs-Yf coup der	0.004563 pm/V3		-0.001253 pm/V3

X offset sens		32.00 nm/V			32.00 nm/V

Fast mag 0		0.90				1.00

Fast mag 1		0.550				0.550

Fast arg		2.50				2.50

Fast arg derate		0.00 1/V			0.00 1/V

Fast cal freq		2.44 Hz			2.44 Hz

Piezo cal		440 V				440 V

Allow rotation		Allow				Allow

Minimum scan rate	0.100 Hz			0.100 Hz

Old				New

Y fast sens		33.28 nm/V			33.67 nm/V

Y fast derate		0.02283nm/V2			0.02558 nm/V2

Y slow sens		36.38 nm/V			37.20 nm/V

Y slow derate		0.02839 nm/V2		0.03184 nm/V2

Ys-Yf coupling	0.05932 nm/V2		0.05196 nm/V2

Ys-Yf coup der	0.05173 pm/V3		-0.006216 pm/V3

Ys-Xf coupling	0.01204 nm/V2		0.01329 nm/V2

Ys-Xf coup der	0.01312 pm/V3		0.01467 pm/V3

Y offset sens		32.00 nm/V			32.00 nm/V

Slow mag 0		1.10				1.10

Slow mag 1		0.600				0.600

Slow arg		2.40				2.50

Slow arg derate	0.00 1/V			0.00 1/V

Slow cal freq		4.77 mHz			4.77 mHz

Rounding 		0.00				0.00

Orthogonality		0.840 				0.00 

Error (x-y direction)	1.7% and 3.2 %		1.2% and 1.4 %