------------------------------

Date:    Wed, 13 Mar 2002 08:27:07 -0800

From:    Clint Callahan <clint@rasy.com>

Subject: Re: noise on contact mode NanoscopeIII afm

Message-ID: <#6>

Shiho

Imaging bare silicon in contact is always fairly difficult. The

surface gets charged just from being scanned and this affects the

force the tip exerts on the sample. Are you able to successfully scan

the sample in contact before doing lithography? If you can scan

before but not after the lithography then the lithography is charging

the surface even more.

You should get an alpha source and have that blasting the sample

while you scan. Nice camera shops carry a 1" brush called a

Staticmaster that has a removable alpha source that perfectly fits a

MultiMode.

If you can't get contact to work using an alpha source, then you will

have to try using a tip that works in both tapping and contact. MFM

tips work well in tapping and they are blunt enough and soft enough

that they work okay in contact. Since they are coated they are

conductive also. You could do your lithography in contact and then

image in tapping.

I also think that it is possible to do the lithography while tapping.

Clint Callahan

Asylum Research

At 9:41 PM -0800 3/8/02, shiho iwanaga wrote:

>Hello,

>

>I am currently facing with noise problem when scanning few nanometer tall

>object on Si wafer.

>

>On Si wafer I have written oxide lines which is 50nm wide and few

>nanometer tall using nanolithography.  I have NanoscopeIII with software

>version 4.43,

>and scanner type E.  When I scanned its image right after creating oxide

>lines, there are noisy signals, as tall as 1 to 2nm, on the Si surface

>which is supposedly very flat.  When the oxide lines created are 1nm tall, it

>is

>completely buried in noise signals.  I know that the surface should be

>flat, with roughness around 0.1 to 0.3nm since this is result I got when I

>tried scanning it using tapping mode.

>

>I suspect that this noise in contact mode operation comes from the fact

>that the trace and retrace lines on oscillo view do not overlap each

>other.  Topography images are close, but they are often 20nm apart.  I

>have tried various combination of PID values but it did not seem to help it.

>

>Has anyone encountered experience such as this? Please tell me how I can

>improve the SNR.

>

>Thank you

>Shiho

>************************************

----------------------------------------------------------------------

Date:    Fri, 22 Feb 2002 10:39:59 +0100

From:    Marcello Campione <marcello.campione@mater.unimib.it>

Subject: electrostatic charges

Message-ID: <#1>

Hi everyone,

I use to work with organic materials grown on non-conductive substrate.

Recently I had some trouble due to long range interaction between the tip

and the sample. Can anyone suggest a way to "sweep off" electrostatic

charges from the sample surface and the tip?

~---------------------------------------------------------------------

Marcello Campione

Dipartimento di Scienza dei Materiali

Universit=E0 degli Studi di Milano Bicocca

V. Cozzi, 53 20125 Milano ITALY

tel: (+39)0264485012

fax: (+39)0264485400

------------------------------

Date:    Fri, 22 Feb 2002 04:03:36 -0600

From:    Dr Raj Lartius <rlartius@novascan.com>

Subject: Re: electrostatic charges

Message-ID: <#2>

Hi Marcello...

A few things that we have used in the past are exposure to alpha particles

(I.e. Americium from a smoke detector or similar from a camera shop),

regulate humidity and/or use a gas passed through an electrostatic

discharge (EDS) device.

Best Wishes,

Raj

~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~

Novascan Technologies, Inc.

Iowa State University Research Park

2501 North Loop Drive

Ames, Iowa 50010  USA

Voice: 515-296-3164

Fax: 515-296-3144

Web:  www.novascan.com

------------------------------

Date:    Fri, 22 Feb 2002 07:31:56 -0500

From:    Don Chernoff At ASM <donc@asmicro.com>

Subject: Re: electrostatic charges

Message-ID: <#3>

Let me add these comments:

~ Be careful when dismantling the radioactive source in a smoke detector!  I

have never done this and I don't know what safety hazards may exist.

~ At ASM, we prefer "StaticMaster" refill cartridges.  The StaticMaster

provides a Po-210 source (beta particles (electrons), stopped by a sheet of

paper or 10 cm of air, I think).  It is considered safe for consumer use in

USA, but a permit may be required in other countries.  We buy from NRD Co.

directly, but this product is also available from supply companies like Ted

Pella or SPI Supplies.

~ another possibility is to use a "ZeroStat" corona discharge device.

I hope this helps.

Don Chernoff

Advanced Surface Microscopy, Inc.    E-Mail:  donc@asmicro.com

6009 KNYGHTON RD.                    Voice:   317-251-1364

INDIANAPOLIS IN 46220                Toll free: 800-374-8557 (in USA)

web:  http://www.asmicro.com         Fax:     317-254-8690

Our preferred web and email addresses are shown above.