Basic techniques for accurate pipetting

-	Always hold your pipette vertical with the tip downwards; this prevents contamination and loss of your sample in the upper parts of the pipette (and protects the pipette).

- 	Drops of solution will stick to the pipette tip. The negative effect on the accuracy will increase with a decreasing volume. Below 100ul wiping the tip after filling becomes important. Try to wipe with a downward sliding motion, being careful not to touch the pipette tip opening.

- 	Recall that the pipette has two stops:

-	The first stop is for picking up the accurate volume and the release should be done by pushing until the second, this is called forward pipetting.

- 	reverse pipetting, which increases accuracy with hygroscopic solutions, is done by picking up by pushing until the second stop and releasing by pushing until the first stop.

-	Make sure all you movements are smooth and slow so that you don't mix air into the volume you are measuring.

- 	When releasing a droplet hold your pipette against the sidewall of the vial to help the release of the total volume from the tip.

October 2005: The pipettes from both 702 and 706 were recently calibrated by Wouter van der

Meer and found to be accurate within 99% for the smallest volumes (10-200 ul) and 99.95% for the 1 ml pipette. One 1 ml Eppendorf pipette was taken for service because of a small error (98.4% accuracy).