The available water qualities in the chemical room of the Huygens Laboratory

Two points of interest:	1	Water quality definitions

2	Properties of our system

As of September 2005:

Properties of our Gradient system from Milli-Q.

- Level of purity is ultra pure grade water.

  - Inorganics 	18 MΩcm-1

  - TOC		<5 ppb

  - Bacteria	< 1CFU/ml

  - Particles	0.2 μm filter

For further definition of the current system look for

(http://www.millipore.com/catalogue.nsf/docs/ZMQS5V0T1)

Before September 2005:

1th The water quality definitions as posed by Elga

- Primary grade water

  - Produced by:	reversed osmosis or weak ion exchange

  - Properties:	quite crude in inorganic, TOC, bacteria and particle content

  - Used for: 	rinsing glassware, preparing reagents

- Deionised water

  - Produced by:	strong ion-exchange

  - Properties:	low ion levels and bacterial content, higher TOC and particle content

  - Used for:	analytical analysis and pharmaceutical solutions

- General laboratory grade water

  - Produced by:	double distillation or multiple other methods.

  - Properties:	low ion, TOC, bacterial and particle content

  - Used for:	general chemistry and qualitative analysis

- Ultra pure grade water

  - Produced by:	“polishing” deionised water

  - Properties:	approaches theoretical levels of pure water

  - Used for:	e.g. HPLC and inductively coupled mass spectrometry

Short table of values:	Primary	Deionised	General grade		Ultra pure

Conductivity (MΩcm-1)	> 1.0		1.0 - 10.0	> 10			> 18.2

TOC			<0.1 ppm			<50 ppb		<5 ppb

2th Properties of our system

- Level of purity is between general and ultra pure grade water.

  - Inorganics 	10-15 MΩcm-1

  - TOC		<50 ppb

  - Bacteria	< 1CFU/ml

  - Particles	0.2 μm filter

    - About the taps:

  - The device uses 2 purification pathways:

    - Raw water  pre treatment  tank

    - Tank  ion & TOC removal  tank

  - The tank and the 2nd pathway have a tap

    - Former gives increasing water quality depending on cycle number

    - Second gives small flow of best quality