# New O- ring from DI, for AFM experiments in liquid

Where is it?	Next to the fluid tip holders in a small petri dish!

How does it work?

First mount a tip in one of the fluid tip holders. Then you can put the O-ring with its small side up in the fluid tip holder (the smaller hole fits tightly around the glass pedestal in the tip holder). Mount a sample in the AFM with just a thin layer of liquid. Insert the tip holder with O-ring and clamp it down. Now you can add more liquid trough one of the small fluid channels on the tip holder. Make sure that you leave the other channel open. The total volume of this fluid cell is around 40 microliters.

How to clean it?

We used SDS and rinsed with water, dried in N2.

The material of the ring is silicone 100% biocompatible.

Problems?

Tip mounting:

Some (small) force with tweezers was needed to stretch the small o-ring hole around the glass pedestal in the tip holder.  Take care not to dislodge/catapult the tip!

Mica/Silicone supports:

The o-ring bottom diameter is ~10mm.  Could there be leakage when clamping down on wet mica with diameter>10mm?  If using smaller (5mm) mica, silicone needs to be cleaned, since it’s enclosed in the liquid space.

Matthew and Dionne